## VM
Requirement No. | Requirement description | Value | Comments from IAS | verification method | Compliant | measured value where applicable - to be filled in for acceptance review | reference document showing the results of the measurement or test - to be filled in for acceptance review | comments from supplier
LW grating
LW grating ruling specification
R-EGA-OPT-2401 | line spacing | Variable |  | D
R-EGA-OPT-2402 | Groove density at vertex | 1740 ± 1 lines/mm |  | T
R-EGA-OPT-2403 | Law groove density | Zemax parameters:  = -7.2998139x10-5± -7.299x10-7 (±1%) 1/T = 1/1.74 –  x (-18.5 ≦ x (mm) ≦ 0) |  | D
R-EGA-OPT-2404 | Groove Alignment | Angle: ± 15’ wrt reference sides Translation: ± 0.5 mm wrt reference side | ± 0.5 mm corresponds to a groove error of ± 0.2 lines/mm | T
R-EGA-OPT-2405 | Groove profile | Blazed | Exact profile to be a proposition by the vendor | T
R-EGA-OPT-2406 | Groove direction | Along y axis |  | T
R-EGA-OPT-2407 | Roughness | RMS roughness increase with respect to the substrates measured roughness < 0.1 nm |  | T
R-EGA-OPT-2408 | Efficiency |  |  | S
R-EGA-OPT-2409 | Ruled area dimensions | Semicircle with 37 mm diameter (2 mm margin). | Centered on OLW (green point in Figure 2) (X, Y) = (0, 0) in the EGA coordinate system (x,y) = (0.000, -2.039) in the LW coordinate system. | T
R-EGA-OPT-2410 | Radius of curvature | The relative change of radius of curvature as a result of the etching process shall be < 1e-5 |  | T
R-EGA-OPT-2411 | 0th order WFE | The delta WFE with respect the substrates measured WFE shall be < 0.2 nm RMS, 1 nm PV | Wrt. best fit conic (R, k). Full aperture measurements scales from to 0.1 to 40 mm | T
Non-optical surfaces requirements
R-EGA-OPT-2301 | Scratch & dig | 40/20 (MIL-O-13830) |  | I
R-EGA-OPT-2302 | Quality of back side reference surface | Polished P3 3/0.25(0.25) < λ / 15 RMS < 0.5 nm rms roughness | At λ = 632.8nm, wrt. best plane Over a >2 cm2 area. For reference. The back surface face is provided polished and should not be modified. | T
R-EGA-OPT-2303 | Quality of other reference surfaces | Polished P1 < λ / 10 RMS | At λ = 632.8nm, wrt best plane Over the entire surface or >2 cm2. Sides B & G in Figure 5 and Figure 8 | T
R-EGA-OPT-2304 | Quality of other sides | grounded |  | I
R-EGA-OPT-2305 | Quality of contact side | Polished | Side A in Figure 5 and Figure 8 | T
R-EGA-OPT-2306 | Perpendicularity of back surface & reference sides | < 15’ (knowledge ± 10”) |  | T
R-EGA-OPT-2307 | Parallelism of reference sides wrt susbstrate reference sides | < 1’ |  | T
R-EGA-OPT-2308 | Orientation of other sides | N/A
R-EGA-OPT-2309 | Angles of contact surface wrt back and reference sides | < 15’ (knowledge ± 10”) |  | T
R-EGA-OPT-2310 | Serial number | Engraved # (LW1, LW2, LW3, …) on a non-reference side, not on the contact side. | Details can be discussed at MRR | I
R-EGA-OPT-2311 | Chamfers | Flat, at 45°, 0.5 mm wide | On all 90° edges | I
tests and metrology
R-VM-1 | verification | The supplier shall give an overview of how the verification, for all methods (Test Inspection Simulation), will be performed and with which accuracy if applicable
Product assurance
R-PA-1 | langage | all documents shall be written in english
R-PA-2 | change request | changes after validation through review (MRR) shall be agreed between IAS team & supplier through a Change Request (CR)
R-PA-3 | non conformities | Final items discrepancy with respect to agreed drawings and requirements shall be traced through NCR (Non-Compliance report)
R-PA-4 | traceability | Each optical element shall be unequivocally identified with a serial number engraved on one non reference surface
R-PA-5 | meetings | monthly progress meeting shall be organized
R-PA-6 | reviews | reviews will be organized as per schedule in section 5.4.2
R-PA-7 | key points | additional key points will be arranged in agreement with both parties
R-PA-8 | handling and cleaning procedure | the supplier shall provide a handling and cleaning (if applicable) procedure for the gratings
R-PA-9 | packing | each optical element shall be packed individually without any contact on the useful optical surface in an ISO4 compatible package