Critères d’acceptation de la carte Interface Board (IB) - Banc de Test IB

| Modifications du Document |  |  |  |  |
| --- | --- | --- | --- | --- |
| Edition | Révision | Date | Pages Modifiées | Observations |
| 0 | 0 | 28/05/2025 | / | Création |

| Reference Documents (RD) |  |  |  |
| --- | --- | --- | --- |
| RD | Title | Reference | Version |

# Prérequis

- La carte IB est entièrement câblée

- Le microcontrôleur est pré-programmé

- La carte IB doit être connectée au banc IB qui assure les fonctions suivantes :

  - Alimentation de la carte avec les tensions +6V, -6V et 3,3V.

  - Simulation de la consommation de la carte HVPA à l’aide de charges.

  - Exécution des tests nécessaires à la vérification des critères d’acceptation.

# Connecteurs

- 7 x Samtec QSE-20-01-L-D-A

- 2 x Tyco 6469076-1

# Critères d’acceptation de la carte IB

## Le test fonctionnel consiste à :

- Vérifier les tensions d’alimentation et la consommation de la carte (tensions, courants, inrush)

- Vérifier le bon fonctionnement du lien SPI

- Vérifier l’acquisition des HouseKeepings

- Tester la ligne reset du microcontrôleur

- Vérifier en temporel le signal de référence de la haute tension REF_HT_DUx pour différentes vitesses de montée en tension et de consignes

Vérifier le ON/OFF de la HT et la sécurité HT

- Vérifier la mise OFF de la HT (après une mise ON)

- Vérifier la mise ON de la HT (après la mise OFF)

- Vérifier la coupure de la HT si détection d’un défaut

- Vérifier la réactivation de la sécurité

- Contrôle de la continuité des 28 pistes des signaux de données HG+, HG-, LG+, LG-

## Tableau récapitulatif des tensions et courants des E/S à vérifier

Remarque :

- x va de 0 à 7

- Les valeurs et tolérances sont à titre indicatif… (28/05/2025)

Alimentations

| Label | Tension (A Injecter et Vérifier) |  | Courant (mA) (A Vérifier) |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Typique (V) | Tolérance | Min | Typique | Max |  |
| +6V | +6 | De +5,9 à +6,3 |  | 460 (AC) |  |  |
| -6V | -6 | De – 6,3 à -5,9 |  | 92 (AC) |  |  |
| +3V3 | +3,3 | +/-5% |  | ~3,5 (AC) |  | Hors communication SPI |
|  | Tension (A Vérifier) |  | Courant (mA) (A Vérifier) La décision d’inclure cette mesure de courant sera prise à l’issue de tests en interne à l’IRAP |  |  | Charges Externes(Ω) ~0,1% |
|  | Typique (V) | Tolérance | Min | Typique | Max |  |
| +6V_DUx | +6 | +/-5% | 2,2 | 2,35 | 2,5 | 2,55k |
| -6V_DUx | -6 | +/-5% | -2,7 | -2,5 | -2,3 | 2,4k |
| +3V3_DUx | +3,3 | +/-5% | 38,2 | 40,2 | 42,3 | 82 |
| +1V8_DUx | +1,8 | +/-5% | 5,5 | 5,8 | 6,2 | 309 |
| V_TRANSFO_DUx | +6V | +/-5% | 4,75 | 5,0 | 5,25 | 1,2k |

Sorties

| Label | Min (V) | Typique (V) | Max (V) | Commentaire | Charges |
| --- | --- | --- | --- | --- | --- |
| V_OSC_DUx | -6 +/-0,3 | - | +6 +/-0,3 | +6V « Haute Tension OFF » -6V « Haute Tension ON » | 1k vers +6V |
| REF_HT_DUx | 0 +/-0,15 | - | 5,0 +/-0,25 | De 0V à 5,0V Consigne de la HT 4,7V est la tension réf max | Oui |
| SPI_MISO | 0 +/-0,15 | - | 3,3 +/-0,175 |  |  |
| Valeurs lues par le microcontrôleur de la carte |  |  |  |  |  |
| Label | Min (mA) | Typique (mA) | Max (V) | Commentaire |  |
| Voltage Measure VMx | TBD | TBD | TBD | Haute Tension 0 < HT < 1500V (TBC) |  |
| Current Measure CMx HK_I_A_x | TBD | TBD | TBD | Courant de la partie HT |  |
| Current Channel CCx HK_I_HVPA_x | 49 TBD | 52 TBD | 55 TBD | Courant de la carte HVPA |  |

Entrées

| Label | Min (V) | Typique (V) | Max (V) | Commentaire |
| --- | --- | --- | --- | --- |
| CRTL_HT_DUx | 0 +/-0,15 | - | 3,2 +/-0,15 | De 0V à 3,2V Contrôle de la HT 3,175V est la tension max HT ON |
| IB_UC_RESET | 0 +/-0,15 | 0 | 3,3 +/-0,175 | 0V Fonctionnement nominal 3,3V Reset du µContrôleur |
| SPIEN_HV | 0 +/-0,15 | - | 3,3 +/-0,175 |  |
| SPI_SCK | 0 +/-0,15 | - | 3,3 +/-0,175 | Fréquence : ~1,5MHz |
| SPI_MOSI | 0 +/-0,15 | - | 3,3 +/-0,175 |  |

Signaux de données

| Label | Continuité en 4 fils |  |  |
| --- | --- | --- | --- |
|  | Min (Ω) | Typique (Ω) | Max (Ω) |
| HG+_DUx, LG+_DUx HG-_DUx, LG-_DUx | - | < 1 | - |

# Détail des fonctions à réaliser (A compléter)

- Vérifier les tensions d’alimentation et la consommation de la carte (tensions, courants, inrush) [EXG#1]

- Vérifier les inrush currents des alimentations : +6V, -6V

- Chronogramme de démarrage des tensions

- Vérifier les tensions en régime nominal

- Vérifier le bon fonctionnement du lien SPI [EXG#2]

- Contrôler les signaux du lien SPI (fréquence, niveau en tension)

- Configurer le microcontrôleur pour passer sur l’application cf. chapitre Lien SPI

- Lire le registre « CURRENT-TRIP » (CT), adresse 0x24, valeur attendue 12489,75mA

- Vérifier l’acquisition des HouseKeepings [EXG#3]

- Lire les valeurs par défaut des registres

- Tester la ligne reset du microcontrôleur [EXG#4]

- Avec HT ON

- Vérifier que la HT se coupe (tension Vosc et registre « disable_hv » (DIS) adresse 0x40)

- Vérifier que le microcontrôleur est à nouveau accessible

- Vérifier le bon fonctionnement de la gestion de la haute tension

- Vérifier en temporel les signaux REF_HT_DUx et CTRL_HT pour différentes vitesses de montée en tension (registre RS) et de consigne (registre VSx), vérifier également la mesure de la HT par le registre Voltage Measure VMx [EXG#5]

- Vérifier le ON/OFF de la HT et la sécurité HT tab. 1, en jouant sur le paramètre THRESHOLD (registre CTx)

| Exigence |  |  | A vérifier |  |
| --- | --- | --- | --- | --- |
|  | HK_I_A Registre CMx | HT ON/OFF registre SWx (SWicth voltage x) | VOSC (V) | Bit DU du registre : Disable_HV (DIS) | COMMENT |
| [EXG#6] | ≤ THRS (Current Trip) | HT ON puis HT OFF | +6 (disable) | 1 | Vérifier la mise OFF de la HT après une mise ON HT on puis off, protection on |
| [EXG#7] | ≤ THRS (Current Trip) | HT OFF puis HT ON | -6 (enable) | 1 | Vérifier la mise ON de la HT après la mise OFF HT off puis on, protection on |
| [EXG#8] | ≤ THRS Puis >THRS (Current Trip) | HT ON puis HK_I_A >THRS | +6 (disable) | 0 | Vérifier la coupure de la HT si détection d’un défaut protection devient off |
| [EXG#9] | ≤ THRS (Current Trip) | Réactivation de la sécurité le registre disable_hv (DIS) puis HT ON | -6 (enable) | 1 | Vérifier la réactivation de la sécurité protection on |

Labels apparaissant dans le schéma électrique :

REF_HT_Dux	: Tension de référence pour le channel x en V

CTRL_HT_Dux	: Tension mesurée pour le channel x en V

HK_I_A_x	: Courant de consommation du CW en µA pour le channel x (conso au niveau du transformateur de la carte HVPA).  Correspond au registre Current Measure DUx

HK_I_HVPA_x	: Mesure du courant de consommation totale de la carte HVPA en µA pour le channel x (interne à la carte IB). Permet de détecter la présence d’un DU

- Contrôle de la continuité des 28 pistes des signaux de données HG+, HG-, LG+, LG- [EXG#10]

- Vérifier la continuité en 4fils

# Lien SPI

Configuration du lien SPI:

- The clock polarity is active on the high state: CPOL = 0

- The clock phase starts at the first edge: CPHA = 1

- The SPI serial data transfers start with the MSB

- Nominal frequency: 1,5MHz

- Voltages: 3,3V

Fonctionnement lors de la mise ON de la carte :

- Le microcontrôleur démarre sur le bootloader

- En attente de commandes pour passer sur l’application

1 - Send READ_STATUS to the bootloader

2 - Check the status value, 30 -> The bootloader is well launched, no command has been launched

3 - Send LOAD_VECTORS to the bootloader

4 – Send JUMP_TO_APPLICATION to the bootloader

- Diagramme de démarrage du microcontrôleur de la carte IB

A la fin de cette séquence, les tests fonctionnels de la carte peuvent être déroulés.

Remarque :

- Toutes les données acquises seront sauvegardées dans une base de données

- Le nombre maximal de cycles de connexion/déconnexion (matage/dématage) des connecteurs doit être respecté

- Les séquences de vecteurs tests seront fournies

# Synoptique général de la carte IB V6 PROTO

# Exemple de banc de test

- Proposition de banc de test

# SPI SLC Mapping Table (IB Application v4.6)

## Global Variables

| Adresse | +0x00 | +0x02 | +0x04 | +0x06 | +0x08 | +0x0A | +0x0C | +0x0E |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0x0000 | MS | MS | KEY | KEY | RS | RS | SRST | SRST |
| 0x0010 | TEMP | TEMP | SN | SN | ID_FPM | ID_FPM | ID_IB | ID_IB |
| 0x0020 | SC | SC | CT | CT | AutoC | AutoC | OFFST | OFFST |
| 0x0030 | VSL | VSL | SWV | SWV | SI2C | SI2C |  |  |
| 0x0040 | DIS | DIS | LSC | LSC | LIC | LIC | LCV | LCV |
| 0x0050 | INP | INP | INM | INM | I2CS | I2CS |  |  |

## Per Channel Variables (CH0 to CH6)

| Adresse | +0x00 | +0x02 | +0x04 | +0x06 | +0x08 | +0x0A | +0x0C | +0x0E |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0x0060 | CS0 | CS0 | SW0 | SW0 | ID_DU0 | ID_DU0 | CT0 | CT0 |
| 0x0070 | VS0 | VS0 | CC0 | CC0 | VM0 | VM0 | CM0 | CM0 |
| 0x0080 | CS1 | CS1 | SW1 | SW1 | ID_DU1 | ID_DU1 | CT1 | CT1 |
| 0x0090 | VS1 | VS1 | CC1 | CC1 | VM1 | VM1 | CM1 | CM1 |
| 0x00A0 | CS2 | CS2 | SW2 | SW2 | ID_DU2 | ID_DU2 | CT2 | CT2 |
| 0x00B0 | VS2 | VS2 | CC2 | CC2 | VM2 | VM2 | CM2 | CM2 |
| 0x00C0 | CS3 | CS3 | SW3 | SW3 | ID_DU3 | ID_DU3 | CT3 | CT3 |
| 0x00D0 | VS3 | VS3 | CC3 | CC3 | VM3 | VM3 | CM3 | CM3 |
| 0x00E0 | CS4 | CS4 | SW4 | SW4 | ID_DU4 | ID_DU4 | CT4 | CT4 |
| 0x00F0 | VS4 | VS4 | CC4 | CC4 | VM4 | VM4 | CM4 | CM4 |
| 0x0100 | CS5 | CS5 | SW5 | SW5 | ID_DU5 | ID_DU5 | CT5 | CT5 |
| 0x0110 | VS5 | VS5 | CC5 | CC5 | VM5 | VM5 | CM5 | CM5 |
| 0x0120 | CS6 | CS6 | SW6 | SW6 | ID_DU6 | ID_DU6 | CT6 | CT6 |
| 0x0130 | VS6 | VS6 | CC6 | CC6 | VM6 | VM6 | CM6 | CM6 |

## Extended Channel Variables (VT, VC, HVm, SC)  v4.6 IB Code

| Adresse | +0x00 | +0x02 | +0x04 | +0x06 | +0x08 | +0x0A | +0x0C | +0x0E |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0x0150 | VT0 | VT0 | VC0 | VC0 | HVm0 | HVm0 | SC0 | SC0 |
| 0x0160 | VT1 | VT1 | VC1 | VC1 | HVm1 | HVm1 | SC1 | SC1 |
| 0x0170 | VT2 | VT2 | VC2 | VC2 | HVm2 | HVm2 | SC2 | SC2 |
| 0x0180 | VT3 | VT3 | VC3 | VC3 | HVm3 | HVm3 | SC3 | SC3 |
| 0x01A0 | VT4 | VT4 | VC4 | VC4 | HVm4 | HVm4 | SC4 | SC4 |
| 0x01B0 | VT5 | VT5 | VC5 | VC5 | HVm5 | HVm5 | SC5 | SC5 |
| 0x01C0 | VT6 | VT6 | VC6 | VC5 | HVm6 | HVm6 | SC6 | SC6 |

| Address | Name | Label | Type | Read/Write | Unit | Comment |
| --- | --- | --- | --- | --- | --- | --- |
| 0x0000 | MS | Module Status | unsigned 32 bits | Read |  | Bit0 : HV_DU0 state Bit1 : HV_DU1 state Bit2 : HV_DU2 state Bit3 : HV_DU3 state Bit4 : HV_DU4 state Bit5 : HV_DU5 state Bit6 : HV_DU6 state |
| 0x0004 | KEY |  |  | Read |  | Read 0xABCDEF IB Application is alive |
| 0x0008 | RS | Ramp Speed | float 32 bits | Read/Write | V/s | High voltage Ramp Speed |
| 0x000C | SRST | Software ReSet |  | Read/Write |  | Write 0xABCDDCBA Application jumps to the Bootloader |
| 0x0010 | TEMP | Temporary variable | float 32 bits | Read |  | Used for debugging |
| 0x0014 | SN | IB soft Serial Number | unsigned 32 bits | Read | ASCII |  |
| 0x0018 | ID_FPM | FPM Identity |  |  |  |  |
| 0x001C | ID_IB | IB Identity |  |  |  |  |
| 0x0020 | SC | Switch Current | int32 |  |  | - 0 : applied Current Trip = 0 for all DUs - 1 : applied Current Trip = CTx for all DUx (i.e. all DUs are set to their own CT value) - 2 : applied Current Trip = max value for all DUs (i.e. the tripping is disable) |
| 0x0024 | CT | Current Trip general | float 32 bits | Read/Write | μA | (THReShold) |
| 0x0028 | AutoC | AutoCalibration | int32 |  |  | autocalibration of VCx variable |
| 0x002C | OFFST |  | int32 |  |  |  |
| 0x0030 | VSL | Voltage Set Low | float 32 bits | Read/Write | V |  |
| 0x0034 | SWV | Mode de fonctionnement de la haute tension Switch Voltage | unsigned 16 bits | Read/Write |  | Modes: 0=Low (<400V) 1=Nominal 2=OFF 3=Star Transit |
| 0x0038 | SI2C | Switch i2c | int32 |  |  | The default value is 0. Set to 1 always run i2c messages to ADC chips on the IB |
| 0x0040 | DIS | Disable HV | unsigned 16 bits | Read |  | Bits 0–6: HV DUx lock Reports the default and allows to reset channels after overcurrent trigger |
| 0x0044 | LSC | Upper Limit Coeff | float 32 bits | Read/Write |  | Upper limit coefficient for the Star transit mode |
| 0x0048 | LIC | Lower Limit Coeff | float 32 bits | Read/Write |  | lower limit coefficient for the Star transit mode |
| 0x004C | LCV | Voltage Control Limit | float 32 bits | Read/Write | V | Control limit for the voltage for the Star transit mode) |
| 0x0050 | INP | Positive Increment | float 32 bits | Read/Write |  | Positive increment for the Star transit mode |
| 0x0054 | INM | Negative Increment | float 32 bits | Read/Write |  | Negative increment for the Star transit mode |
| 0x0058 | I2CS | I2C State | int32 | Read |  | Indicates the active state of the I2C, such as an LED blink status |
| 0x0060 | CS0 | Channel Status DU0 | unsigned 16 bits | Read/Write |  | 0: DU not detected 1: DU detected |
| 0x0064 | SW0 | Switch Voltage DU0 | unsigned 16 bits | Read/Write |  | Modes 0–3 |
| 0x0068 | ID_DU0 | Identity DU0 |  |  |  |  |
| 0x006C | CT0 | Current Trip DU0 | float 32 bits | Read/Write | μA | Current trip of the DU |
| 0x0070 | VS0 | Voltage Set DU0 | float 32 bits | Read/Write | V | Voltage set of the DU |
| 0x0074 | CC0 | Current HVPA DU0 | float 32 bits | Read | V |  |
| 0x0078 | VM0 | Voltage Measure DU0 | float 32 bits | Read | V | HT Value of the DU |
| 0x007C | CM0 | Current Measure DU0 | float 32 bits | Read | μA | Consumption current of the high voltage system of the DU |
| 0x0080 | CS1 | Channel Status DU1 | unsigned 16 bits | Read/Write |  | 0: DU not detected 1: DU detected |
| 0x0084 | SW1 | Switch Voltage DU1 | unsigned 16 bits | Read/Write |  | Modes 0–3 |
| 0x0088 | ID_DU1 | Identity DU1 |  |  |  |  |
| 0x008C | CT1 | Current Trip DU1 | float 32 bits | Read/Write | μA | Current trip of the DU |
| 0x0090 | VS1 | Voltage Set DU1 | float 32 bits | Read/Write | V | Voltage set of the DU |
| 0x0094 | CC1 | Current HVPA DU1 | float 32 bits | Read | V |  |
| 0x0098 | VM1 | Voltage Measure DU1 | float 32 bits | Read | V | HT Value of the DU |
| 0x009C | CM1 | Current Measure DU1 | float 32 bits | Read | μA | Consumption current of the high voltage system of the DU |
| 0x00A0 | CS2 | Channel Status DU2 | unsigned 16 bits | Read/Write |  | 0: DU not detected 1: DU detected |
| 0x00A4 | SW2 | Switch Voltage DU2 | unsigned 16 bits | Read/Write |  | Modes 0–3 |
| 0x00A8 | ID_DU2 | Identity DU2 |  |  |  |  |
| 0x00AC | CT2 | Current Trip DU2 | float 32 bits | Read/Write | μA | Current trip of the DU |
| 0x00B0 | VS2 | Voltage Set DU2 | float 32 bits | Read/Write | V | Voltage set of the DU |
| 0x00B4 | CC2 | Current HVPA DU2 | float 32 bits | Read | V |  |
| 0x00B8 | VM2 | Voltage Measure DU2 | float 32 bits | Read | V | HT Value of the DU |
| 0x00BC | CM2 | Current Measure DU2 | float 32 bits | Read | μA | Consumption current of the high voltage system of the DU |
| 0x00C0 | CS3 | Channel Status DU3 | unsigned 16 bits | Read/Write |  | 0: DU not detected 1: DU detected |
| 0x00C4 | SW3 | Switch Voltage DU3 | unsigned 16 bits | Read/Write |  | Modes 0–3 |
| 0x00C8 | ID_DU3 | Identity DU3 |  |  |  |  |
| 0x00CC | CT3 | Current Trip DU3 | float 32 bits | Read/Write | μA | Current trip of the DU |
| 0x00D0 | VS3 | Voltage Set DU3 | float 32 bits | Read/Write | V | Voltage set of the DU |
| 0x00D4 | CC3 | Current HVPA DU3 | float 32 bits | Read | V |  |
| 0x00D8 | VM3 | Voltage Measure DU3 | float 32 bits | Read | V | HT Value of the DU |
| 0x00DC | CM3 | Current Measure DU3 | float 32 bits | Read | μA | Consumption current of the high voltage system of the DU |
| 0x00E0 | CS4 | Channel Status DU4 | unsigned 16 bits | Read/Write |  | 0: DU not detected 1: DU detected |
| 0x00E4 | SW4 | Switch Voltage DU4 | unsigned 16 bits | Read/Write |  | Modes 0–3 |
| 0x00E8 | ID_DU4 | Identity DU4 |  |  |  |  |
| 0x00EC | CT4 | Current Trip DU4 | float 32 bits | Read/Write | μA | Current trip of the DU |
| 0x00F0 | VS4 | Voltage Set DU4 | float 32 bits | Read/Write | V | Voltage set of the DU |
| 0x00F4 | CC4 | Current HVPA DU4 | float 32 bits | Read | V |  |
| 0x00F8 | VM4 | Voltage Measure DU4 | float 32 bits | Read | V | HT Value of the DU |
| 0x00FC | CM4 | Current Measure DU4 | float 32 bits | Read | μA | Consumption current of the high voltage system of the DU |
| 0x0100 | CS5 | Channel Status DU5 | unsigned 16 bits | Read/Write |  | 0: DU not detected 1: DU detected |
| 0x0104 | SW5 | Switch Voltage DU5 | unsigned 16 bits | Read/Write |  | Modes 0–3 |
| 0x0108 | ID_DU5 | Identity DU5 |  |  |  |  |
| 0x010C | CT5 | Current Trip DU5 | float 32 bits | Read/Write | μA | Current trip of the DU |
| 0x0110 | VS5 | Voltage Set DU5 | float 32 bits | Read/Write | V | Voltage set of the DU |
| 0x0114 | CC5 | Current HVPA DU5 | float 32 bits | Read | V |  |
| 0x0118 | VM5 | Voltage Measure DU5 | float 32 bits | Read | V | HT Value of the DU |
| 0x011C | CM5 | Current Measure DU5 | float 32 bits | Read | μA | Consumption current of the high voltage system of the DU |
| 0x0120 | CS6 | Channel Status DU6 | unsigned 16 bits | Read/Write |  | 0: DU not detected 1: DU detected |
| 0x0124 | SW6 | Switch Voltage DU6 | unsigned 16 bits | Read/Write |  | Modes 0–3 |
| 0x0128 | ID_DU6 | Identity DU6 |  |  |  |  |
| 0x012C | CT6 | Current Trip DU6 | float 32 bits | Read/Write | μA | Current trip of the DU |
| 0x0130 | VS6 | Voltage Set DU6 | float 32 bits | Read/Write | V | Voltage set of the DU |
| 0x0134 | CC6 | Current HVPA DU6 | float 32 bits | Read | V |  |
| 0x0138 | VM6 | Voltage Measure DU6 | float 32 bits | Read | V | HT Value of the DU |
| 0x013C | CM6 | Current Measure DU6 | float 32 bits | Read | μA | Consumption current of the high voltage system of the DU |
| Debug address (v4.6 du code IB) |  |  |  |  |  |  |
| 0x0150 | VT0 | Voltage Target0 | float 32 bits |  | V |  |
| 0x0154 | VC0 | Voltage Coefficient 0 |  |  |  | internal dev purposes |
| 0x0158 | HVm0 | Mean High Voltage 0 | int32 | Read | mV | in ~10 sec time |
| 0x015C | SC0 | Switch Current 0 | int32 | Read | - | • 0 : Current Trip is set to 0. Can be used to switch security system for all DUs • 1 : Current Trip is set to the nominal value (cf. CT and CTx). Usual value |
| 0x0160 | VT1 | Voltage Target1 | float 32 bits |  | V |  |
| 0x0164 | VC1 | Voltage Coefficient 1 | float 32 bits | Read | V |  |
| 0x0168 | HVm1 | Mean High Voltage 1 | int32 | Read | mV | in ~10 sec time |
| 0x016C | SC1 | Switch Current 1 | int32 | Read | - | Cf SC0 |
| 0x0170 | VT2 | Voltage Target2 | float 32 bits |  | V |  |
| 0x0174 | VC2 | Voltage Coefficient 2 |  |  |  | internal dev purposes |
| 0x0178 | HVm2 | Mean High Voltage 2 | int32 | Read | mV | in ~10 sec time |
| 0x017C | SC2 | Switch Current 2 | int32 | Read | - | Cf SC0 |
| 0x0180 | VT3 | Voltage Target3 | float 32 bits |  | V |  |
| 0x0184 | VC3 | Voltage Coefficient 3 |  |  |  | internal dev purposes |
| 0x0188 | HVm3 | Mean High Voltage 3 | int32 | Read | mV | in ~10 sec time |
| 0x018C | SC3 | Switch Current 3 | int32 | Read | - | Cf SC0 |
| 0x01A0 | VT4 | Voltage Target4 | float 32 bits |  | V |  |
| 0x01A4 | VC4 | Voltage Coefficient 4 |  |  |  | internal dev purposes |
| 0x01A8 | HVm4 | Mean High Voltage 4 | int32 | Read | mV | in ~10 sec time |
| 0x01AC | SC4 | Switch Current 4 | int32 | Read | - | Cf SC0 |
| 0x01B0 | VT5 | Voltage Target5 | float 32 bits |  | V |  |
| 0x01B4 | VC5 | Voltage Coefficient 5 |  |  |  | internal dev purposes |
| 0x01B8 | HVm5 | Mean High Voltage 5 | int32 | Read | mV | in ~10 sec time |
| 0x01BC | SC5 | Switch Current 5 | int32 | Read | - | Cf SC0 |
| 0x01C0 | VT6 | Voltage Target6 | float 32 bits |  | V |  |
| 0x01C4 | VC6 | Voltage Coefficient 6 |  |  |  | internal dev purposes |
| 0x01C8 | HVm6 | Mean High Voltage 6 | int32 | Read | mV | in ~10 sec time |
| 0x01CC | SC6 | Switch Current 6 | int32 | Read | - | Cf SC0 |