Reference : Consultation n°202609 Technical Response Framework (in French = Cadre de Réponse Technique, CRT)

Laser desorption-ionisation source

The candidate indicates in this document the services that he undertakes to implement within the framework of the execution of the contract. Each element is evaluated based on the answers provided in this response framework and all documents or technical data sheets attached by the candidate allowing verification of the offer's compliance with the requirements of the CNRS/DR18 CCP/

Candidates are advised that if the information marked with an asterisk (*) in the CRT is not provided, the offer will be considered incomplete and therefore irregular. It will then not be ranked.

Regarding the other fields, the absence of a response will be considered as a lack of commitment from the candidate on the item concerned and will penalize the score given to the candidate's offer.

The candidate is asked to answer the following questions as precisely as possible:

I – Checklist of Required Technical Specifications, the CCTP in French (items not weighted in the score)

In this section I and its subsections 1-8, the questions are designed to ensure that candidates fully understand the Required Technical Specifications in the CCTP. A negative response to one or more questions below (in I-1 to I-8) could call into question the compatibility and/or the performances of the solution and lead to the elimination of the offer due to irregularities.

- (*) Is the proposed solution compatible with the Orbitrap™ Eclipse IC Ready mass spectrometer available at PhLAM? In particular, once coupled to the OrbitrapTM Eclipse, does it allow:

- to continue operating the combined ESI, APCI and APPI ionization source OptamaxTM NG?

- to continue using the OptamaxTM NG ion source for the automatic calibration of the OrbitrapTM Eclipse spectrometer?

- to use the Easy-ICTM internal calibration source of the spectrometer?

- to be disconnected from the spectrometer for maintenance and cleaning of the mass spectrometer?

- to transmit the ions generated by laser desorption-ionization for analysis by the mass spectrometer?

- to trigger the acquisition of data (full-scan, tandem MS) by the mass spectrometer in synchronization with the laser desorption-ionization and the transfer of the ions thus generated to the OrbitrapTM Eclipse?

- to be coupled to the OrbitrapTM Eclipse spectrometer as installed on an optical table at PhLAM?

h) to operate the ion optics and preserve the vacuum quality in the Orbitrap™ mass spectrometer?

- (*) Specifics of the laser desorption-ionization system proposed by the candidate, desorption section:

Is a laser of wavelength 349 or 355 nm integrated to the solution to perform desorption-ionization?

Does the solution allow for laser desorption-ionization with a laser beam in “reflected” mode (in incidence from above the sample surface)?

Is it possible to use the desorption laser focused through an objective located below the sample (“transmission” mode)?

Would it be possible, at the time of installation of the proposed solution, to integrate the 1064 nm laser (model Crylas-DSS1064-450) from PhLAM and configure it for laser desorption in "transmission" as well as "reflected" modes?

- (*) Specifics of the laser desorption-ionization system proposed by the candidate, post-ionization section:

- Is a 266 nm laser integrated to the solution in order to perform post-ionization, and synchronized with the desorption laser so that each desorption laser pulse is emitted alternately with each post-ionization laser pulse?

Indicate if the optimal alignment and focusing of the post-ionization laser system, in order to increase the ionization of desorbed molecules, will be performed during installation. Will users be trained to perform such adjustments during post-installation training?

- (*) Visualization of the sample in the desorption-ionization source:

  - Does the system allow imaging of the sample in reflected light using a visible light source located above the sample, and a first dedicated camera?

  - Does the system allow imaging of the sample in transmitted light, using: the same light source located above the sample, objectives used to focus the laser of the desorption source in "transmission" mode, and a second dedicated camera?

- (*) Environment of the samples in the desorption-ionization source:

  - In case pumping is performed, is it carried out with an oil-free pump (dry pump)?

  - Is a valve used to isolate the proposed system from the vacuum of the mass spectrometer during the loading/unloading of the sample into the laser desorption-ionization source?

  - Is it possible to inject gases (at least, N2) into the desorption-ionization source using a leak valve during operation of the desorption-ionization source? Is this compatible with the operation of the desorption-ionization source, the transfer of ions to the mass spectrometer, and the acquisition of mass spectra by the spectrometer?

- (*) Sample positioning in the desorption-ionization source:

Do the mechanical design and the sample-positioning stages of the proposed complete solution allow‒in both “transmitted” and “reflected” modes‒analyses in laser desorption-ionization as well as visualization of the sample, over a zone of at least 40 x 25 mm?

- (*) Transfer and injection of ions (generated by laser desorption-ionization and/or laser post-ionization) into the mass spectrometer:

  - Can the proposed solution transfer to the mass spectrometer both the positive and negative ions generated by laser desorption-ionization as well as by post-ionization?

  - Does the proposed solution allow to reach the optimal mass resolution of 1 million (in m/Δm FWHM at m/z 200 u.) in mass spectra acquired for ions generated by laser desorption-ionization (and/or laser post-ionization) and by the ESI/APPI/APCI source?

- (*) Control systems. Controls of the ion source (including triggering of the desorption and post-ionization lasers, sample positioning, sample visualization) and of the ion transfer system, and their coupling with the mass spectrometer should be provided that may operate on a single computer. Are features in the following list included in the proposed system?

- connectivity with the OrbitrapTM Eclipse mass spectrometer allowing for full scan mass spectrometry and multistage mass spectrometry (MSn, following HCD, CID and/or UVPD dissociations) on ions generated by the laser desorption-ionization events (on single spots and for automated acquisition of sample-raster mass spectrometric imaging datasets) as well as on ions from the ESI/APCI/APPI source?

- readout and control of the parameters relevant to the transfer and injection in the mass spectrometer of the ions generated by the laser desorption-(post-)ionization events and the ESI/APCI/APPI source?

- control of the synchronization of the laser desorption with the ion transfer system and mass spectrometer data acquisition?

- control of the gate/slide valve protecting the vacuum quality in the mass spectrometer during sample exchange?

- readout of the pressure(s) in the sample environment and ion transfer system (if different)?

- readout of the desorption laser energies and readout of the post-ionization laser energy?

- control of pulse rates for the desorption laser and for the post-ionization laser?

- control of the number of laser shots at each pixel of single-spot analyses or of mass spectrometric imaging?

- control of the delays between desorption and post-ionization laser pulses?

- switching on/off the light source used for sample visualization?

- sample visualization on the transmitted- and reflected-light cameras, and recording of images from each camera?

- control of sample stage positions, speeds and accelerations?

- control of the sample positioning for “bulk” mode of analysis with the possibility to analyze multiple spots of the sample surface with displacement of the XY stages during the injection time of the OrbitrapTM Eclipse mass spectrometer?

- acquisition of mass spectrometric imaging data, using protocols that comprise means of programming X-Y sample translation steps for automated acquisition of mass spectrometry imaging datasets over user-selected areas of the sample?

II – Technical value criterion (worth 56% of the overall score)

- (*) Specifics of the laser desorption-ionization system proposed by the candidate, desorption section (total 9 %):

  - Indicate the maximal energy per laser pulse (µJ) of the desorption laser integrated into the system.

  - Indicate the usable range of repetition rates (Hz) of the desorption laser integrated into the system.

  - Indicate the beam quality (M² index) of the desorption laser integrated into the system.

  - What is the angle of incidence (in degrees) of the desorption laser in "reflected" mode (incidence above the sample and its sample holder)?

  - Indicate the dimensions, in micrometers, of the desorption laser beam once focused on the sample in "reflected" mode, and the possibilities of adjusting these dimensions by translation of an optical element.

  - Indicate the dimensions, in micrometers, of the desorption laser beam once focused through a 1 mm glass slide in "transmission" mode and the possibilities of adjusting these dimensions by motorized translation of the focusing objective.

  - Concerning the optical elements enabling laser desorption-ionization, when the damage threshold is known, precise it as an average number of laser pulses that a given element can withstand at the highest known (by the supplier) energy.

- (*) Specifics of the laser desorption-ionization system proposed by the candidate, post-ionization section (total 9 %):

  - Indicate the maximal energy per laser pulse (µJ) of the laser integrated into the system to perform post-ionization of the neutrals produced by the desorption laser.

  - Indicate the usable range of repetition rates (Hz) of this post-ionization laser.

  - Indicate the beam size (at the laser output) of this post-ionization laser.

  - Indicate the orientation of the post-ionization laser relative to the surface of the sample.

  - Describe the optical path clearance (in millimeters above the sample surface) in the system accepting the passage of post-ionization laser beams between the sample and the ion transmission system to the mass spectrometer.

  - Describe the method used to synchronize the post-ionization laser shots with the desorption laser shots (shots emitted alternately with a delay that can be adjusted by the users).

- (*) Visualization of the sample in the desorption-ionization source (7 %):

  - What type of illumination is used for visualizing the samples in transmitted and reflected light? If available: what is its luminous intensity (in lumens)?

  - What is the lateral resolution of reflected-mode imaging (in µm) at the sample surface, and the number of pixels of the camera dedicated to this mode?

  - What is the lateral resolution of transmitted-mode imaging (in µm) at the sample surface, and the number of pixels of the camera dedicated to this mode?

  - Concerning the optical elements enabling these visualizations of the sample, when the damage threshold is known, precise it as an average number of laser pulses that a given element can withstand at the highest known (by the supplier) energy.

- (*) Environment of the samples in the desorption-ionization source (7 %):

  - What is the pressure inside the laser desorption-ionization source? What are the minimum and maximum accepted pressure values? Do those pressure depend on the filling gas?

  - Is pumping carried out with an oil-free pump (dry pump)?

  - What gases (N2, Ar, He, Xe, air, or other), injected into the desorption-ionization source using a leak valve, are compatible with the operation of the desorption-ionization source, the transfer of ions to the mass spectrometer, and the acquisition of mass spectra by the spectrometer?

  - Indicate the maximal dimensions of the samples that can be used in the laser desorption-ionization source (lengths, widths, and thicknesses, in millimeters). Can samples deposited onto glass slides measuring 75x25x1 mm (L*W*thickness), metallized with indium-tin oxide (ITO), be used?

- (*) Sample positioning (6 %). Describe the properties of the sample-positioning stages provided for X/Y travel in the plane parallel to the sample surface, including:

- the total number of stages provided,

- their maximum travel distances (in mm),

- the type(s) of motor(s) (piezoelectric, mechanical) used,

- their backlash (reproducibility of displacement to a given position for coarse pluri-millimetric travel, in µm),

- the response time between a displacement command and the beginning of the displacement (in milliseconds),

- The travel speeds (in mm/s) and resolutions of displacement (in nm).

- (*) Transfer and injection of ions (generated by laser desorption-ionization and/or laser post-ionization) into the mass spectrometer (9%):

  - Describe the type of system used to transfer ions generated by laser desorption-ionization (and post-ionization) into the mass spectrometer.

  - Indicate the mass range of the ions (in mass/charge ratios m/z) that can be transferred and injected into the mass spectrometer by the proposed system.

  - Estimate the efficiency (in %) of the transfer and injection of the ions generated in the desorption ionization source to the mass spectrometer, over the mass/charge ratio range m/z of 50-1000 u.

  - How can the solution help with the removal of neutral particles/molecules produced by laser desorption-ionization and by the ESI/APPI/APCI sources, in order to limit their injection into the ion optics of the mass spectrometer?

- (*) Control systems (9%).

- Describe how the analysis in laser desorption-ionization is triggered by the control system, at the level of 1) the laser desorption-ionization and the laser post-ionization events, 2) the transfer of ions into the mass spectrometer, and 3) the start of the ion analysis by the mass spectrometer.

Describe the protocols allowing for the acquisition of mass spectrometric imaging data, including means of programming X-Y sample translation steps for automated acquisition of mass spectrometry imaging datasets over user-selected areas of the sample.

Describe mass spectrometry imaging software that can be used to process the mass spectrometry imaging data generated with the proposed instrumental solution.

Is a dedicated computer included in the offer to carry the controls associated with the proposed solution? What is its operating system?

III – Price criterion (worth 30% of the overall score)

- Price

(*) The candidate details the total cost of the proposed solution:

IV – Criterion on delivery time and timing of the training (worth 4% of the overall score)

- (*) The candidate specifies the time required for delivery to the PhLAM laboratory (from the date of notification of the contract, this time cannot exceed 18 weeks) (2%):

- (*) The candidate specifies, from the date of delivery, the timeframe for installing the instrument, including validation tests and user training? (This timeframe may not exceed 1 month) (2%):

V – Criterion on after-sales services (worth 6 % of the overall score)

- (*) Response time to a request (email/phone, in working day(s)), and estimated on-site intervention delay if a repair is requested (in working day(s).  (2 %)

- (*) Specify the proposed warranty period (minimum 2 year): (2%)

- (*) Mandatory PSE Description of preventive and/or corrective maintenance contracts. This description may include: the number and type of annual preventive maintenance, and the number of spare parts included in the optional maintenance plan. (2%)

VI – Environmental criterion (worth 4% of the overall score)

(*) System repairability.

Estimation of the durability of the most wear-sensitive components associated with the proposed instrumental solution (example: vacuum pumps, translation stages, objectives under IR/UV laser irradiation). Availability period for spare parts (from the start date of the warranty).