# ANNEX 2 – TECHNICAL SPECIFICATIONS

## 1. General Technical and Compliance Requirements

Unless otherwise stated under a specific item, all offered items shall meet the following minimum requirements:

- All items shall be new, unused, unused in previous service, and of current production.

- All items shall be of medical grade and suitable for the intended operational use.

- Where applicable, offered equipment shall comply with recognised international standards and regulatory requirements, including CE Mark, FDA clearance/approval, ISO certification, or equivalent internationally recognised standards, as applicable to the offered product category.

- The bidder shall submit a fully detailed technical offer for all offered items, including:

  - manufacturer name;

  - exact model / catalogue number;

  - country origin: the item shall comply with applicable European standards (EN) or equivalent internationally recognized standards and shall be approved or accepted by the Royal Medical Services (RMS) for use by the Jordan Armed Forces (JAF). Documentary evidence of compliance shall be provided with the technical offer.

  - country where the manufacturer is based;

  - delivery time;

  - warranty period;

  - full technical description/specifications;

  - all accessories and options included in the offer.

- The bidder shall submit clear original technical brochures / catalogues / datasheets for all offered items.

- The bidder shall submit compliance sheets in tabular format against the technical specifications, with a clear Yes / No response for each requirement and reference to the relevant brochure page and line number. Offers not complying with this requirement may be rejected.

- All required accessories and consumables necessary for safe, complete and functional operation of the offered equipment shall be clearly identified. Any accessories and consumables shall be manufacturer-approved and, where required, priced separately.

- The bidder shall demonstrate capacity for after-sales support, installation where applicable, and maintenance support through authorized / certified biomedical service engineers with relevant experience.

- Warranty documentation shall be provided for all applicable equipment. Unless otherwise stated in the contract, supplied equipment shall be covered by a minimum 12-month warranty.

- All supplied items shall be fully compatible with the requirements of the Jordan Armed Forces (JAF) and the Royal Medical Services (RMS) and shall remain subject to technical verification prior to final acceptance.

- Where applicable, power requirements shall be suitable for local conditions and clearly stated in the offer. RMS general requirements refer to single phase 220V / 50Hz or 3-phase 380V, where relevant to the supplied system.

- Where applicable, pre-installation and site-readiness requirements necessary for installation and operation of the offered equipment shall be clearly identified by the bidder. RMS general conditions place responsibility for pre-installation works on the supplier where applicable.

The technical specifications are based on the requirements provided by the Jordan Armed Forces (JAF) / Royal Medical Services (RMS). Where only functional requirements are defined, bidders shall provide detailed technical characteristics demonstrating compliance.

# 2. Item-Specific Minimum Technical Requirements

## 2.1 Emergency Trolley (Crash Cart), Complete

Quantity: 1 unit

Minimum requirements:

- Heavy-duty, advanced mobile design with easy access to components;

- Polymer or equivalent durable construction;

- Full-swivel castors with brakes, minimum 10 cm diameter;

- Ergonomic handle;

- Quick-release master locking / security latch mechanism with plastic seals;

- All-round bumpers;

- Minimum 5 drawers with dividers, including small, medium and large drawers;

- Must include at least:

  - defibrillator / monitor shelf;

  - oxygen cylinder holder with regulator;

  - sharps container with holder;

  - waste bin with holder;

  - backboard;

  - one IV pole.

## 2.2 Patient Monitor (ECG / NIBP / SpO2)

Quantity: 2 units

Minimum requirements:

- Reputable manufacturer; durable, compact, heavy-duty design;

- Preconfigured or compact modular monitor;

- Operates on mains AC supply without external converters/adapters;

- Suitable for adult, paediatric and neonatal use;

- Ability to measure/display at least:

  - ECG waveform;

  - NIBP;

  - heart rate;

  - body temperature;

  - SpO2;

- Original accessories, including sensors / cuffs / probes as applicable;

- Mounted on original trolley with accessories basket;

- Real-time waveform and numeric display;

- Colour screen of approximately 12.1 inches, resolution not less than 800 x 600;

- Battery backup minimum 90 minutes;

- Escalating audio-visual alarm that cannot be permanently silenced;

- Minimum 72-hour trend capacity;

- Ability to display 4 waveforms simultaneously.

## 2.3 Defibrillator (AED / Manual, Adult & Pediatric)

Quantity: 1 unit

Minimum requirements:

- Reputable manufacturer; durable, compact, heavy-duty design;

- Both manual and AED defibrillation capability;

- Biphasic waveform technology;

- ECG monitoring capability;

- Synchronisation capability;

- Lightweight portable design;

- Colour LCD display;

- Adult and paediatric capability;

- Rechargeable battery and charger;

- Complete with all accessories required for operation.

## 2.4 Examination Couch

Quantity: 1 unit

Minimum requirements:

- Fixed examination couch;

- Heavy-duty design and construction;

- Stainless steel or epoxy-coated steel frame;

- Approximate overall dimensions (H x W x L): 85 x 60 x 200 cm;

- The surface shall be at least 5 cm thick, padded and upholstered with wear-resistant and disinfectant-tolerant material. The upholstery shall be well padded, durable, easy to clean and disinfect, and made from PVC or similar material;Adjustable back rest with secure locking mechanism.

## 2.5 Hospital Bed, Manual (Short-Term Observation)

Quantity: 4 units

Minimum requirements:

- Heavy-duty construction suitable for rugged use;

- Bed weight ≥ 140 kg (bed only, without accessories and mattress);

- Safe working load to be at least 200 kg;

- The dimensions of the sleeping surface platform shall be a minimum of 90 x 200 cm;

- Flat sleep surface;Four-section platform configuration, including:

  - Backrest (head of bed) section;

  - Seat section;

  - Thigh section;

  - Foot (calf section);

- The bed shall be manually operated and shall include the following manoeuvres:

  - Height adjustment of the bed: ≤ 50 cm to ≥ 75 cm, measured without mattress from floor to the top of the platform;

  - Knee gatch (break): 0° to ≥ 20°;

  - Trendelenburg / reverse Trendelenburg: +12° / -12° or more;

- Zero clearance between mattress and headboard;

- Clearance between dropped-down side rails and the floor in the bed’s lowest position (under-bed clearance) shall be at least 10 cm;

- Easy-release manual CPR, fully accessible even when side rails are in the lowest position;

- Backrest (head of bed) angle display;

- Left and right side rails shall:

  - Be split and heavy-duty model;

  - Offer zero transfer gap;

  - Be manually raised and dropped down with damping and top locking mechanism;

  - Include a safe gap concept preventing patient fall and entrapment;

  - Cover at least 2/3 of the sleeping platform;

- Heavy-duty non-metallic headboard and footboard (at least the headboard shall be removable);

- Four heavy-duty, medical-grade, electro-conductive, non-marking castors of not less than 5-inch diameter;

- Central brake system with steering function;

- Four original protective revolving bumpers;

- Four IV-pole sockets;

- One compatible IV pole with double hooks;

- Drainage bag hook located on each side of the bed;

- Sleeping surfaces, quantity 4, with the following minimum requirements:

  - Pressure ulcer preventive;

  - Foam core;

  - Patient weight capacity up to 150 kg or more;

  - Fire retardant, anti-microbial and anti-fungus;

  - Complete with non-latex, fluid-resistant and washable cover.

## 2.6 Suction Machine, Medical Grade, Mobile

Quantity: 2 units

Minimum requirements:

- Durable, heavy-duty construction, reliable, ergonomic design and easy to use;

- Mobile, mounted on antistatic castors, at least 4, with a sturdy stable base that resists tipping over;

- On/off switch with light indicator;

- Maximum flow rate not less than 30 L/min;

- Easy control of vacuum level through a knob or a switch on the control panel;

- Maximum vacuum not less than 675 mmHg;

- Display of vacuum through a gauge that is easy to read, with high visibility mmHg and kPa graduation;

- Two original reusable and autoclavable jars to be included with the unit, with lids providing secure seal and float-type cut-off valve to prevent overflow. All jar sizes to be offered;

- Bacteria filter to minimize cross infection, available from different manufacturers;

- Silicone transparent tubes;

- Splash-proof design;

- Oil-free, maintenance-free unit;

- Quiet pump operation, below 45 dB;

- Two extra original jars to be provided with the unit.

## 2.7 Medication Cart

Quantity: 2 units

Minimum requirements:

- Heavy-duty design;

- The medication cart shall be a unit-dose cart with individual bins or drawers for each patient’s dosage;

- Ergonomic heavy-duty design;

- Constructed from polymer material;

- Shall include an upper work surface;

- At least 30 large storage space bins of approximate size 20 cm x 10 cm x 14 cm (L x W x H);

- Large lower utility drawer with removable adjustable dividers, including 3 longitudinal sections;

- Master lock or locking mechanism securing the entire cart, which must be unlocked before any individual compartment can be opened;

- Push handle;

- All-round or corner bumpers;

- Waste container;

- Mobile cart mounted on large swivel castors with locks.

## 2.8 Portable Operation Lamp

Quantity: 1 unit

Minimum requirements:

- Surgical light system based on high-performance LED light technology and from a reputable well-known company;

- Single light head: compact, ergonomically shaped;

- Mounting type: floor type;

- Continuously adjustable focalization;

- Colour temperature within the range of 3,500 K – 5,000 K;

- Colour Rendering Index (CRI): up to 95 or better;

- Deep saturated red colour index (R9): must be specified;

- The light shall achieve a central illumination of at least 140,000 Lux at 1 m distance;

- Shadow management or shadow-less;

- Heavy-duty original mobile stand;

- Autoclavable handles, 3 pairs;

- LED life: 50,000 hours or more;

- Intensity illumination control;

- Top side of light head with smooth surfaces and even transitions for easy and quick disinfection/cleaning.

## 2.9 Electrosurgical Unit

Quantity: 1 unit

Minimum Requirements:

- From a reputable well-known brand;

- Application: suitable for a wide range of major surgeries including at least:

  - General Surgery;

  - Orthopaedic Surgery;

  - Laparoscopic Surgery;

  - Caesarean Section Surgery;

  - ENT Surgery;

- Controls to include:

  - Monopolar mode;

  - Bipolar mode;

  - Pure Cut;

  - Blend Cut;

  - Coagulation;

  - Spray Coagulation;

- Automatic stable output control to maintain selected power level across different tissue types;

- Neutral patient plate monitoring system with indicator;

- Maximum Cut Output Power: 300–400 W;

- Maximum Coagulation Output Power: 80 W or better;

- Leakage Current: less than 20 mA;

- Operating panel with touch buttons for mode selection and power adjustment;

- Digital LCD display showing power levels and error codes;

- Monopolar handle connection: International 3-pin plug type;

- Handle control and footswitch control capability;

- Audible and visual alarm system with self-test capability;

- Original heavy-duty mobile trolley with large twin anti-static lockable castors;

- Safety Type: CF Defibrillation Proof.

Accessories to be quoted separately:

- Adult disposable split patient plate (Qty. 20);

- Bipolar forceps with cable;

- Reusable monopolar pencil;

- Monopolar handle tips:

  - Ball;

  - Loop;

  - Needle;

  - Spatula;

  - Curved spatula;
(Qty. 1 each)

- Hook electrode with cable (Qty. 1);

- Twin spill-proof footswitch;

- All additional cutting and coagulating tips and forceps required for the specified surgical applications.

Additional Requirements:

- Handle tips, forceps and cables must be autoclavable.

### 2.10 Blood Chemistry Analyzer

Quantity: 1 unit

Minimum Requirements:

- Fully automated dry chemistry analyzer suitable for routine biochemical tests;

- Capable of processing a minimum of 5 samples simultaneously;

- STAT sample prioritization capability;

- Sample types:

  - Serum;

  - Plasma;

  - Whole Blood;

- Colour touch screen display;

- Refrigerated reagent storage;

- Integrated printer or external printer;

- Complete with all accessories required for normal operation;

- Complete with starter kit, calibrators, controls and consumables required for commissioning and operation;

- Suitable for operation on 220V / 50Hz power supply.

### 2.11 Mobile Digital X-Ray System with Wireless Flat Panel Detector

Quantity: 1 unit

Minimum Requirements:

A heavy-duty digital mobile radiographic unit is required for radiographic imaging of patients in various hospital settings.

The unit shall satisfy the following minimum requirements:

- X-ray tube and flat panel detector shall comply with RMS Special Terms requirements;

- Built-in detector battery charger;

- One extra detector battery included;

- Full microprocessor control of all functions with self-test routine;

- Controlled mobility via motion system;

- High battery storage capacity;

- Input charging voltage: 100–240 VAC;

- System shall operate on battery power and while connected to AC power;

- Safety anti-collision system or safety bumpers;

- Touch monitor;

- Tube arm maximum reach: 110 cm;

- System start-up time shall be specified by the supplier;

- Collimator with integrated variable filter;

- Manually adjustable multi-leaf collimator rotatable ±90°;

- LED collimator light;

- Maximum SID (Source to Image Distance): 190 cm;

- Minimum SID (Source to Image Distance): 65 cm;

### X-Ray Generator

- Effective nominal power: 30 kW or higher;

- mAs range: 0.5 – 320 mAs or higher;

- kVp range: 50 – 125 kVp or higher;

### Detector

- Wireless portable detector;

- Rigid design;

- Operates on both 2.4 GHz and 5 GHz frequency bands;

- Pixel pitch: 150 μm or less;

- Scintillator type: CsI;

- Preview image time: 8 seconds or less;

- Active detector area: 35 × 43 cm or larger;

- Detector weight not exceeding 4.0 kg including battery;

- Anti-scatter grid included;

### X-Ray Tube

- Anode heat capacity: 120 KHU;

- Tube rotation: ±90°;

- Dual focal spot:

  - Small focal spot ≤ 0.7 mm;

  - Large focal spot ≤ 1.3 mm;

OR

  - Nominal focal spot ≤ 0.8 mm;

- Anode rotation speed: minimum 2700 rpm approximately;

### Connectivity

- Full DICOM 3.0 compatibility;

- RIS compatibility;

- HIS compatibility;

- PACS compatibility.

### Optional Item

- Dose Area Product (DAP) meter to be quoted separately.

### 2.12 Endodontic Rotary Motor Unit

Quantity: 1 unit

Minimum Requirements:

- Durable and heavy-duty unit from a reputable well-known brand;

- Main purpose: Root Canal Preparation;

- The unit shall include at least the following:

a. Control Unit

  - Compact lightweight control unit;

  - LCD display;

  - Capable of controlling speed and torque of the micro-motor;

b. Electric Micro-Motor

  - Cordless type;

  - Hand-control operation;

c. Contra-Angle Handpiece

  - One contra-angle handpiece;

  - Gear reduction ratio: 16:1;

  - Quick connect/disconnect type;

  - Push-button chuck mechanism;

d. Charger

  - Handpiece charger;

  - Wireless charging type (without cable);

- Auto-reverse function;

- All handpieces and heads shall be autoclavable at 134°C;

- Operated by built-in rechargeable batteries.

### 2.13 Haematology Analyzer (3-Part Differential)

Quantity: 1 unit

Minimum Requirements:

- Latest and highest-end model within the required product range;

- Automated haematology analyser with minimum 3-part differential capability;

- Utilizes whole blood samples;

- Capable of measuring not less than 20 parameters, including at least:

  - WBC;

  - RBC;

  - Hb;

  - PCV;

  - MCV;

  - MCH;

  - MCHC;

  - Platelets;

- Histograms for:

  - WBC;

  - RBC;

  - Platelets;

- Histograms shall be displayed on the screen and printed with results;

- Throughput: not less than 55 samples/hour;

- Cyanide-free reagents for haemoglobin and WBC measurements;

- Sample volume not exceeding 100 µL;

- Touch screen display;

- Integrated barcode reader;

- Minimum three quality control files:

  - Low;

  - Normal;

  - High;

- Capability to display Levey-Jennings graphs on screen;

- Built-in thermal printer;

- Memory capacity not less than 1,000 samples, with or without histograms;

- Complete start-up kit including all reagents, consumables, piercer, electrodes and maintenance kits;

- Start-up kit shall be sufficient for 10,000 samples and include at least one level of daily quality control.

### 2.14 Ordinary Light Microscope

Quantity: 1 unit

Minimum Requirements:

- 5-fold revolving objective nosepiece, rear facing;

- Semi-Apochromatic infinity objectives:

  - 4x;

  - 10x;

  - 20x;

  - 40x;

  - 100x;

- All objectives shall be of the same class and suitable for use with up to 22 mm field of view;

- Ergonomic binocular tube with 22 mm field of view;

- Pair of adjustable eyepieces:

  - 10x / 22 mm;

  - With interpupillary distance (IPD) adjustment;

- Transmitted light illumination with collector and Halogen lamp not less than 50 W; or LED illumination;

- Coaxial double-knob focusing system with adjustable height stop;

- Swing-out condenser with aperture iris diaphragm and variable Koehler illumination;

- Mechanical stage with scanning area not less than 75 x 50 mm;

- Slide holder suitable for up to two slides;

- Simple polarizing set included;

- Standard filter set included;

- Adjustable light intensity control.

### 2.15 Urine Analyzer

Quantity: 1 unit

Minimum Requirements:

- Latest and highest-end model within the required product range;

- Uses urine test strips;

- Built-in thermal printer or external printer;

- Throughput not less than approximately 60 strips/hour;

- Capable of measuring at least the following parameters:

  - Glucose;

  - Protein;

  - Bilirubin;

  - Urobilinogen;

  - pH;

  - Specific Gravity;

  - Ketone Bodies;

  - Occult Blood (RBCs);

  - Leukocytes (WBCs);

- Data storage capacity for approximately 100 patient results or more;

- Capability to connect to:

  - Hospital Information System (HIS);

  - Laboratory Information System (LIS).

## 2.16 Dental Unit

Quantity: 1 unit

Minimum requirements:

- Heavy-duty dental unit from a reputable manufacturer;

### A. Dental Chair

- Heavy-duty, robust unit from a well-known reputable manufacturer;

- All chair movements shall be electrically motorized;

- Heavy-duty maintenance-free unit;

- Microprocessor-controlled chair with standard safety features, including movement limitation when obstacles are present;

- Seamless and replaceable upholstery with high-grade hygienic finish;

- Arm rests on both sides;

  - Right arm rest shall be swivel type to facilitate patient entry and exit;

- Extendable jointed headrest;

- Minimum lifting capacity: 170 kg;

- Chair movements shall include at least:

  - Vertical seat adjustment;

  - Backrest adjustment;

  - Trendelenburg position;

  - Return-to-zero position;

  - Programmable treatment positions.

### B. Dentist's Element

- Mounted on an articulated delivery arm with drop-down configuration;

- Equipped with locking brake system;

- Capable of accommodating at least five instrument positions, including:

a. High-Speed Modules

  - Two high-speed handpiece modules;

b. Electric Micro-Motor Module

  - Brushless type;

  - Same manufacturer/brand as the dental unit;

  - Speed range up to 40,000 rpm;

c. Three-Way Syringe

  - Built-in module;

  - Round-edged ergonomic design;

  - Two removable autoclavable nozzles;

d. Piezo Scaler Module

  - Built-in piezo scaler;

  - Complete with four (4) tapered triangular tips;

- All hoses shall be top-hanging to avoid entanglement with patients and staff;

- Autoclavable removable swivelling tray;

  - Minimum size: 30 cm × 20 cm;

- One wired multifunctional foot control capable of controlling:

  - Chair movements;

  - Handpiece operation;

- Touch control panel capable of controlling:

  - All chair movements;

  - Dental light;

  - Cup filling;

  - Spittoon rinsing.

### C. Assistant's Element

- Chair-mounted spittoon movable with chair vertical movements;

- All major controls and filters shall be located inside the unit for easy maintenance and service access;

- Assistant element shall include suction system controls and accessories;

- Flexible, smooth and easy-to-clean/disinfect suction hoses;

- LED Light Cure Handpiece:

  - Power intensity: 2,000 mW/cm² or higher;

  - Wavelength: approximately 430–490 nm or better;

- Touch control panel capable of controlling:

  - All chair movements;

  - Cup filling;

  - Spittoon rinsing;

- Dry vacuum separation system:

  - Automatic operation;

  - Located in the hydro-group;

- Dental unit shall be equipped with central suction system connections;

- Spittoon bowl shall:

  - Be made of ceramic or equivalent material;

  - Be acid and alkaline resistant;

  - Be easy to clean;

  - Be swivel type;

  - Include rinsing system;

  - Include adjustable time control;

  - Include water flow regulator.

### D. Dental Light

- Easy to clean and compact design;

- LED type;

- Light intensity not less than 50,000 Lux;

- LED service life not less than 30,000 operating hours.

### E. Water Container

- A manually refillable water container/reservoir shall be provided;

- The reservoir shall be capable of supplying water to the dental handpieces as a backup water source.

### F. Dentist's and Assistant's Stools

- One stool for the dentist;

- One stool for the assistant;

- Mounted on five castors;

- Adjustable height by gas spring mechanism

### G. Turbine Handpiece

Quantity: 5 handpieces with each dental unit

- From reputable manufacturers such as Bien-Air, NSK, W&H or Sirona;

- Quick connect/disconnect type;

- Autoclavable up to 134°C;

- Speed: 320,000 rpm or higher;

- Power: 18 W or higher;

- Internal cooling mechanism (Internal Irrigation);

- Push-button bur locking mechanism;

- Ceramic ball bearing;

- Quick coupling (4-hole type) to be quoted separately (Qty. 2).

### H. Low Speed Contra-Angle Handpiece

Quantity: 5 handpieces with each dental unit

- Bien-Air, NSK, W&H or Sirona;

- Electrically driven handpiece;

- Contra-angle type;

- Blue ring;

- 1:1 transmission ratio;

- Internal cooling mechanism;

- Universal E-type connection;

- Push-button bur locking mechanism.

### I. Low Speed Straight Handpiece

Quantity: 1 handpiece with each dental unit

- Bien-Air, NSK, W&H or Sirona;

- Electrically driven handpiece;

- Straight type;

- Blue ring;

- 1:1 transmission ratio;

- Internal cooling mechanism;

- Universal E-type connection;

- Twist tension chuck system.

### 2.17 Dental Cabinet Trolley

Quantity: 3 units

Minimum Requirements:

- Worktop material shall be Corian or equivalent;

- Approximate worktop dimensions:

  - Depth: 55 cm;

  - Height: 90 cm;

- Each cabinet shall include at least:

a. Drawer Units

  - Quantity: 2 units;

  - Each unit shall include:

    - Minimum three medium-size pull-out drawers;

    - One deep storage drawer;

    - Necessary partitions and trays.

b. Double-Door Cupboard Unit

  - Quantity: 1 unit;

  - Double hinged doors;

  - Two adjustable shelves.

c. Single-Door Cupboard Unit

  - Quantity: 1 unit;

  - One hinged door;

  - Three adjustable shelves.

- Four heavy-duty swivel castors.

### 2.18 Dental Digital X-Ray

Quantity: 1 unit

Minimum Requirements:

- Heavy-duty, robust unit from a well-known reputable manufacturer;

- Wall-mounted type;

- Generator type:

  - High-frequency constant potential technology (DC type);

- Focal spot:

  - 0.4 mm or smaller;

- X-ray tube voltage:

  - 60 kV or higher;

- X-ray tube current:

  - Greater than 4 mA;

- Incrementally adjustable exposure time;

- Pre-programmed exposure settings for adults and paediatric patients;

- Pre-programmed exposure settings for multiple tooth positions;

- Focus-to-skin distance:

  - 20 cm;

- Optional longer focus-to-skin distance to be quoted separately;

- Hand switch:

  - Push-button type;

  - Small size;

  - Cord length minimum 3 meters;

  - Wireless hand switch acceptable;

- Arm length:

  - Minimum 80 cm.

### 2.19 Mobile X-Ray Table with Integrated Detector Bucky Tray

Quantity: 1 unit

Minimum Requirements:

- Mobile patient table with integrated detector Bucky tray suitable for general radiography examinations;

- Heavy-duty design suitable for use with mobile X-ray units;

- Lightweight and stable structure with anti-rust coating;

- Detector Bucky tray integrated underneath the tabletop;

- Detector Bucky tray shall accommodate different detector sizes, including at least Agfa Detector Model XD 14;

- Grid shall be removable and insertable into the Bucky tray;

- Radiolucent tabletop with approximate dimensions:

  - Length: 202 cm (±2 cm);

  - Width: 70 cm (±2 cm);

- Distance from tabletop to floor:

  - 70 cm (±2 cm);

- Maximum tabletop load:

  - 135 kg (297 lbs);

- Easy-to-clean antibacterial surface;

- Four swivel castors with foot brakes (minimum two locking brakes).

### 2.20 Class B Autoclave Sterilizer

Quantity: 2 units

Minimum Requirements:

- Class B Autoclave;

- The autoclave shall comply with EN 13060 for small autoclaves;

- The unit shall include at least the following cycles:

a. Class B Cycle (134°C)

  - For wrapped instruments and porous materials;

b. Class B Cycle (121°C)

  - For wrapped instruments and porous materials;

c. Vacuum Leak Test Cycle

d. Bowie-Dick Test Cycle

- Time for Class B typical load universal program at 134°C shall not exceed 50 minutes, including drying phase;

- Manual water feed (stand-alone unit), with no requirement for external water connection;

- Built-in steam generator;

  - Units with internal heater design shall be rejected;

- Built-in water quality sensor (conductivity sensor);

  - Cycle shall not start in case of unacceptable water quality;

- Chamber constructed from high-quality 316 stainless steel;

- High-performance membrane vacuum pump;

- Door lock system preventing door opening during sterilization cycle;

- LCD screen displaying:

  - Program;

  - Cycle phase;

  - Remaining time;

  - Error messages;

- Minimum three original trays and tray holder;

- Integrated feed-water tank and waste-water tank;

  - Minimum capacity: 3 litres each;

  - Units with external tanks shall be rejected;

- External dimensions suitable for installation on a standard bench with 60 cm depth;

- Power rating:

  - Less than 12 Amperes;

- Water distiller complete with storage bottle.

## 2.21 Airway Set (OPA + NPA, All Sizes)

Quantity: 40 sets

Minimum requirements:

- Complete range of adult and paediatric sizes;

- Medical-grade material;

- Individually packed, clean / sterile as applicable;

- Suitable for emergency airway management.

## 2.22 McGill Forceps (Adult & Pediatric)

Quantity: 4 units

Minimum requirements:

- Stainless steel;

- Reusable and autoclavable;

- Suitable adult and paediatric sizes;

- Medical-grade finish.

## 2.23 Laryngoscope Set + ET Tubes (Assorted Sizes)

Quantity: 2 sets

Minimum requirements:

- Complete laryngoscope set with handle and blades;

- Assorted adult and paediatric sizes as applicable;

- Endotracheal tubes included in assorted sizes;

- Medical-grade materials;

- Reusable / autoclavable components where applicable.

## 2.24 Bag Valve Mask (BVM), Adult & Pediatric

Quantity: 5 units

Minimum requirements:

- Self-inflating manual resuscitation bag;

- Adult and paediatric sizes;

- Complete with masks and oxygen reservoir where applicable;

- Suitable for emergency ventilation use.

## 2.25 Nasal Cannula, Adult & Pediatric

Quantity: 250 units

Minimum requirements:

- Medical-grade disposable oxygen cannula;

- Adult and paediatric sizes as required;

- Suitable for oxygen therapy use.

## 2.26 Non-Rebreather Oxygen Masks (3 Sizes)

Quantity: 200 units

Minimum requirements:

- Disposable medical-grade oxygen masks;

- Complete with reservoir bag;

- Three sizes as required;

- Suitable for emergency oxygen administration.

## 2.27 IV Start Packs

Quantity: 500 units

Minimum requirements:

- Sterile, single-use IV initiation packs;

- Complete kits suitable for peripheral IV insertion;

- Medical-grade consumables.

## 2.28 Normal Saline 0.9%, 1,000 ml IV Bags

Quantity: 1,000 units

Minimum requirements:

- Sterile intravenous normal saline 0.9%;

- Volume: 1,000 ml;

- Medical-grade packaging suitable for clinical use.

## 2.29 Suction Machine Consumables (1-Year Supply)

Quantity: 1 set

Minimum requirements:

- Fully compatible with the supplied suction machine;

- Includes tubing, filters, collection accessories and other routine operating consumables as applicable;

- Sufficient for one-year operational use under normal service conditions.

## 2.30 Electrosurgical Unit Consumables (1-Year Supply)

Quantity: 1 set

Minimum requirements:

- Fully compatible with supplied electrosurgical unit;

- Includes patient plates, electrodes, tips and routine operating consumables;

- Sufficient for one-year operation.

## 2.31 Blood Chemistry Analyzer Reagents and Consumables (1-Year Supply)

Quantity: 1 set

Minimum requirements:

- Fully compatible with supplied analyzer;

- Includes reagents, calibrators, controls and consumables;

- Sufficient for one-year operation.

## 2.32 Digital Mobile X-Ray Consumables and Accessories (1-Year Supply)

Quantity: 1 set

Minimum requirements:

- Fully compatible with supplied digital X-ray system;

- Includes accessories and consumables required for routine operation;

- Sufficient for one-year operation.

## 2.33 Endodontic Consumables (1-Year Supply)

Quantity: 1 set

Minimum requirements:

- Fully compatible with supplied endodontic motor;

- Includes files, accessories and routine consumables;

- Sufficient for one-year operation.

## 2.34 Defibrillator Consumables (Pads + Battery, 1-Year Supply)

Quantity: 1 set

Minimum requirements:

- Adult and paediatric pads;

- Spare battery;

- Fully compatible with the supplied defibrillator;

- Sufficient for one-year operational use under normal service conditions.

## 3. Mandatory Compliance Statement

All supplied items must fully comply with these technical specifications and with the relevant requirements of the Jordan Armed Forces (JAF) / Royal Medical Services (RMS).

The bidder shall clearly demonstrate compliance through:

- technical brochures / catalogues;

- detailed technical datasheets;

- completed compliance sheets;

- exact references to brochure page and line numbers.

Any deviation shall be explicitly stated in the offer. Unstated deviation, incomplete compliance documentation, or supply of non-compliant items may constitute grounds for rejection.